2011-11-22 at 16:41

Logged in as 9WGC8P1\Administrator

Instrument: Triple TOF 5600

Model #: 1032150/N

Serial #: AY21451105

Instrument Optimization Ver: 3.4172.40

Instrument performance before tuning...

TOF High Resolution

| Mass (Da) | Found At (Da) | Height (cps) | Area | Resolution | Error (ppm) |
| --- | --- | --- | --- | --- | --- |
| 132.9049 | 132.9061 | 6.72E+03 | 2.57E+04 | 24 604 | 8.8 |
| 829.5393 | 829.5483 | 1.14E+04 | 6.89E+04 | 37 665 | 10.8 |

Product Ion High Resolution

| Mass (Da) | Found At (Da) | Height (cps) | Area | Resolution | Error (ppm) |
| --- | --- | --- | --- | --- | --- |
| 185.1285 | 185.0745 | 1.88E+04 | 7.78E+04 | 26 901 | 291.5 |
| 215.1390 | 215.0763 | 1.35E+04 | 5.40E+04 | 29 224 | 291.4 |
| 298.2125 | 298.1267 | 4.79E+04 | 2.08E+05 | 32 288 | 287.9 |
| 381.2496 | 381.1396 | 5.33E+04 | 2.29E+05 | 36 524 | 288.5 |
| 494.3337 | 494.1909 | 7.59E+04 | 3.28E+05 | 41 428 | 289.0 |
| 607.4178 | 607.2421 | 6.00E+04 | 2.89E+05 | 40 605 | 289.2 |
| 625.4283 | 625.2471 | 3.32E+04 | 1.70E+05 | 38 585 | 289.7 |
| 724.4967 | 724.2865 | 2.47E+04 | 1.33E+05 | 39 736 | 290.2 |
| 811.5288 | 811.2929 | 3.42E+03 | 2.04E+04 | 37 818 | 290.7 |

Tuning instrument...

Resolution optimization started

Original Values

RO2: 27.00

VS1: -0.30

HST: 0.00

FOR: 7.34

OFS: -23.03

MGV: -974.05

Adjusting ITC to: 0.9

RO2: 26.82

Adjusting ITC to: 1.8

VS1: -0.31

HST: 0.00

OFS: -37.15

MGV: -973.92

FOR: 8.86

MGV: -974.10

Final Values

RO2: 26.82

VS1: -0.31

HST: 0.00

FOR: 8.86

OFS: -37.15

MGV: -974.10

Orignal Aligment Values

Chan 0 Alignment Positive0.075381,-26.723546

Chan 2 Alignment Positive-0.115324,2.172406

Chan 3 Alignment Positive-0.186673,-7.213924

Chan 0 Alignment Negative-0.005767,-26.742754

Chan 2 Alignment Negative-0.022015,1.392134

Chan 3 Alignment Negative-0.031232,-8.062721

Aligning channels

Final Aligment Values

Chan 0 Alignment Positive0.009783,-28.728461

Chan 2 Alignment Positive-0.049191,-0.207410

Chan 3 Alignment Positive-0.086566,-8.987333

Chan 0 Alignment Negative-0.005767,-26.742754

Chan 2 Alignment Negative-0.022015,1.392134

Chan 3 Alignment Negative-0.031232,-8.062721

TOF Calibration started

TOF Calibration started

Instrument performance after tuning...

TOF High Resolution

| Mass (Da) | Found At (Da) | Height (cps) | Area | Resolution | Error (ppm) |
| --- | --- | --- | --- | --- | --- |
| 132.9049 | 132.9045 | 7.32E+03 | 2.75E+04 | 24 767 | 2.6 |
| 829.5393 | 829.5381 | 1.26E+04 | 7.10E+04 | 40 950 | 1.5 |

Product Ion High Resolution

| Mass (Da) | Found At (Da) | Height (cps) | Area | Resolution | Error (ppm) |
| --- | --- | --- | --- | --- | --- |
| 185.1285 | 185.1281 | 2.07E+04 | 8.03E+04 | 28 418 | 1.9 |
| 215.1390 | 215.1385 | 1.47E+04 | 5.67E+04 | 30 622 | 2.3 |
| 298.2125 | 298.2128 | 5.39E+04 | 2.14E+05 | 35 084 | 0.9 |
| 381.2496 | 381.2503 | 5.55E+04 | 2.36E+05 | 37 221 | 1.9 |
| 494.3337 | 494.3339 | 8.39E+04 | 3.34E+05 | 44 717 | 0.3 |
| 607.4178 | 607.4180 | 6.69E+04 | 2.99E+05 | 43 929 | 0.3 |
| 625.4283 | 625.4283 | 3.76E+04 | 1.82E+05 | 41 397 | 0.1 |
| 724.4967 | 724.4964 | 2.84E+04 | 1.44E+05 | 42 183 | 0.4 |
| 811.5288 | 811.5276 | 4.07E+03 | 2.29E+04 | 40 751 | 1.4 |

**************

Instrument Optimization has modified the instrument settings with the best values it has found. If you are not satisfied with the performance of the instrument, you can restore the settings to a previous state by using the 'File / Restore Instrument Settings' command.

**************

Done