2011-11-23 at 08:55

Logged in as 9WGC8P1\Administrator

Instrument: Triple TOF 5600

Model #: 1032150/N

Serial #: AY21451105

Instrument Optimization Ver: 3.4172.40

Instrument performance before tuning...

TOF High Resolution

| Mass (Da) | Found At (Da) | Height (cps) | Area | Resolution | Error (ppm) |
| --- | --- | --- | --- | --- | --- |
| 132.9049 | 132.9105 | 9.42E+03 | 3.56E+04 | 24 794 | 42.1 |
| 829.5393 | 829.5740 | 1.36E+04 | 7.66E+04 | 40 766 | 41.8 |

Product Ion High Resolution

| Mass (Da) | Found At (Da) | Height (cps) | Area | Resolution | Error (ppm) |
| --- | --- | --- | --- | --- | --- |
| 185.1285 | 185.1361 | 2.19E+04 | 8.66E+04 | 28 455 | 41.3 |
| 215.1390 | 215.1479 | 1.51E+04 | 6.07E+04 | 29 247 | 41.5 |
| 298.2125 | 298.2255 | 5.64E+04 | 2.30E+05 | 33 738 | 43.6 |
| 381.2496 | 381.2665 | 5.71E+04 | 2.50E+05 | 35 925 | 44.4 |
| 494.3337 | 494.3551 | 8.86E+04 | 3.46E+05 | 45 240 | 43.2 |
| 607.4178 | 607.4437 | 7.35E+04 | 3.16E+05 | 45 455 | 42.7 |
| 625.4283 | 625.4551 | 4.07E+04 | 2.00E+05 | 41 022 | 42.9 |
| 724.4967 | 724.5275 | 3.13E+04 | 1.59E+05 | 42 217 | 42.5 |
| 811.5288 | 811.5627 | 4.61E+03 | 2.54E+04 | 40 799 | 41.8 |

Tuning instrument...

TOF Detector optimization started

AAO_AAOEventSampleChangeStatus

Acquisition was aborted.

TOF Detector optimization cancelled