2011-11-23 at 10:53

Logged in as 9WGC8P1\Administrator

Instrument: Triple TOF 5600

Model #: 1032150/N

Serial #: AY21451105

Instrument Optimization Ver: 3.4172.40

Instrument performance before tuning...

Enhance One Ion

| Mass (Da) | Mode | Height (cps) | Area |
| --- | --- | --- | --- |
| 381.2496 | Normal | 5.27E+03 | 4.16E+04 |
| 381.2496 | Enhanced | 1.26E+04 | 1.06E+05 |

Tuning instrument...

original IRD coef. = 3.49

original IRW coef. = 1.35

TOF Enhance All optimization started

Original Enhancing parameter values

RO2= 23.1

IQ3= 25.4

IRP= 20.4

IRD: 3.49

IRW: 1.35

Optimal IRD coef. = 3.49

Optimal IRW coef. = 1.35

IQ3 increase: 5.0

IRP decrease: 1.7

RO2 decrease: 3.7

Final IRD coef. = 3.5

Final IRW coef. = 1.4

Final Enhancing parameter values

RO2= 23.1

IQ3= 25.4

IRP= 20.4

TOF Enhance All optimization finished

Instrument performance after tuning...

Enhance One Ion

| Mass (Da) | Mode | Height (cps) | Area |
| --- | --- | --- | --- |
| 381.2496 | Normal | 5.41E+03 | 4.19E+04 |
| 381.2496 | Enhanced | 1.27E+04 | 1.07E+05 |

**************

Instrument Optimization has modified the instrument settings with the best values it has found. If you are not satisfied with the performance of the instrument, you can restore the settings to a previous state by using the 'File / Restore Instrument Settings' command.

**************

Done