2011-11-29 at 11:31

Logged in as 9WGC8P1\Administrator

Instrument: Triple TOF 5600

Model #: 1032150/N

Serial #: AY21451105

Instrument Optimization Ver: 3.4172.40

Instrument performance before tuning...

Q1

| Mass (Da) | Found At (Da) | Height (cps) | Width | Error (Da) |
| --- | --- | --- | --- | --- |
| 59.05 | 59.05 | 2.68E+04 | 0.72 | 0.0 |
| 175.13 | 175.14 | 4.16E+03 | 0.74 | 0.0 |
| 442.34 | 442.34 | 7.91E+04 | 0.66 | 0.0 |
| 674.50 | 674.50 | 4.66E+04 | 0.65 | 0.0 |
| 906.67 | 906.68 | 1.87E+05 | 0.66 | 0.0 |
| 1196.88 | 1196.88 | 4.00E+04 | 0.67 | 0.0 |

Tuning instrument...

Q1 Scan tuning started.

59.0 measuring offset vs. width

59.0 measuring offset vs. width

59.0 measuring offset vs. width

59.0 measuring offset vs. width

59.0 measuring offset vs. width

59.0 measuring offset vs. width

59.0 measuring offset vs. width

59.0 measuring offset vs. width

59.0 measuring offset vs. width

59.0  original resolution offset= 62.335

59.0 optimized resolution offset= 62.305

175.1 measuring offset vs. width

175.1 measuring offset vs. width

175.1 measuring offset vs. width

175.1 measuring offset vs. width

175.1 measuring offset vs. width

175.1 measuring offset vs. width

175.1 measuring offset vs. width

175.1  original resolution offset= 62.510

175.1 optimized resolution offset= 62.491

442.3 measuring offset vs. width

442.3 measuring offset vs. width

442.3 measuring offset vs. width

442.3 measuring offset vs. width

442.3 measuring offset vs. width

442.3  original resolution offset= 62.555

442.3 optimized resolution offset= 62.543

674.5 measuring offset vs. width

674.5 measuring offset vs. width

674.5 measuring offset vs. width

674.5 measuring offset vs. width

674.5  original resolution offset= 62.557

674.5 optimized resolution offset= 62.548

906.7 measuring offset vs. width

906.7 measuring offset vs. width

906.7 measuring offset vs. width

906.7 measuring offset vs. width

906.7 measuring offset vs. width

906.7  original resolution offset= 62.553

906.7 optimized resolution offset= 62.550

1196.9 measuring offset vs. width

1196.9 measuring offset vs. width

1196.9 measuring offset vs. width

1196.9 measuring offset vs. width

1196.9 measuring offset vs. width

1196.9 measuring offset vs. width

1196.9 measuring offset vs. width

1196.9 measuring offset vs. width

1196.9  original resolution offset= 62.555

1196.9 optimized resolution offset= 62.552

tuning finished

Q1 Scan Calibration started. Dwell time =	10.00

59.04914 m/z found at 59.04 m/z

59.04914 m/z    original DAC value =	 2443

59.04914 m/z  calibrated DAC value =	 2443

175.1 low signal

442.3374 m/z found at 442.29 m/z

442.3374 m/z    original DAC value =	 18498

442.3374 m/z  calibrated DAC value =	 18496

674.50484 m/z found at 674.47 m/z

674.50484 m/z    original DAC value =	 28225

674.50484 m/z  calibrated DAC value =	 28223

906.67228 m/z found at 906.66 m/z

906.67228 m/z    original DAC value =	 37951

906.67228 m/z  calibrated DAC value =	 37951

1196.88158 m/z found at 1196.86 m/z

1196.88158 m/z    original DAC value =	 50111

1196.88158 m/z  calibrated DAC value =	 50110

Calibration finished

Instrument performance after tuning...

Q1

| Mass (Da) | Found At (Da) | Height (cps) | Width | Error (Da) |
| --- | --- | --- | --- | --- |
| 59.05 | 59.05 | 2.76E+04 | 0.74 | 0.0 |
| 175.13 | 175.12 | 4.38E+03 | 0.77 | 0.0 |
| 442.34 | 442.35 | 8.56E+04 | 0.76 | 0.0 |
| 674.50 | 674.50 | 5.20E+04 | 0.76 | 0.0 |
| 906.67 | 906.65 | 1.93E+05 | 0.73 | 0.0 |
| 1196.88 | 1196.87 | 4.24E+04 | 0.73 | 0.0 |

**************

Instrument Optimization has modified the instrument settings with the best values it has found. If you are not satisfied with the performance of the instrument, you can restore the settings to a previous state by using the 'File / Restore Instrument Settings' command.

**************

Done