2012-03-20 at 13:22

Logged in as 9WGC8P1\Administrator

Instrument: Triple TOF 5600

Model #: 1032150/N

Serial #: AY21451105

Instrument Optimization Ver: 3.4172.40

Instrument performance before tuning...

TOF High Resolution

| Mass (Da) | Found At (Da) | Height (cps) | Area | Resolution | Error (ppm) |
| --- | --- | --- | --- | --- | --- |
| 146.1176 | 146.1497 | 1.28E+04 | 4.45E+04 | 27 140 | 219.9 |
| 266.1598 | 266.2185 | 1.09E+04 | 4.42E+04 | 32 329 | 220.6 |
| 315.1623 | 315.2318 | 1.00E+04 | 4.06E+04 | 34 624 | 220.8 |
| 354.2122 | 354.2906 | 6.24E+03 | 2.71E+04 | 34 497 | 221.1 |
| 442.2647 | 442.3623 | 7.68E+03 | 3.39E+04 | 38 844 | 220.8 |
| 609.2807 | 609.4149 | 4.84E+03 | 2.44E+04 | 39 282 | 220.3 |
| 618.3695 | 618.5059 | 8.99E+03 | 4.36E+04 | 41 068 | 220.5 |

Product Ion High Resolution

| Mass (Da) | Found At (Da) | Height (cps) | Area | Resolution | Error (ppm) |
| --- | --- | --- | --- | --- | --- |
| 174.0913 | 174.0951 | 4.54E+04 | 1.59E+05 | 29 619 | 21.8 |
| 195.0652 | 195.0698 | 6.85E+04 | 2.47E+05 | 31 463 | 23.4 |
| 236.1281 | 236.1334 | 1.40E+04 | 5.46E+04 | 31 896 | 22.2 |
| 365.1860 | 365.1945 | 1.64E+04 | 6.97E+04 | 36 170 | 23.3 |
| 397.2122 | 397.2215 | 4.91E+04 | 1.96E+05 | 40 108 | 23.5 |
| 448.1966 | 448.2072 | 2.61E+04 | 1.16E+05 | 38 081 | 23.6 |
| 609.2807 | 609.2948 | 4.10E+04 | 1.85E+05 | 43 820 | 23.2 |

Product Ion High Sensitivity

| Mass (Da) | Found At (Da) | Height (cps) | Area | Resolution | Error (ppm) |
| --- | --- | --- | --- | --- | --- |
| 174.0913 | 174.0910 | 7.16E+04 | 3.62E+05 | 20 503 | 1.7 |
| 195.0652 | 195.0646 | 8.84E+04 | 4.15E+05 | 23 403 | 3.0 |
| 236.1281 | 236.1277 | 2.67E+04 | 1.85E+05 | 17 689 | 1.5 |
| 365.1860 | 365.1858 | 2.81E+04 | 2.25E+05 | 18 779 | 0.5 |
| 397.2122 | 397.2116 | 5.81E+04 | 3.91E+05 | 23 415 | 1.6 |
| 448.1966 | 448.1963 | 3.89E+04 | 3.14E+05 | 20 624 | 0.6 |
| 609.2807 | 609.2794 | 4.63E+04 | 3.90E+05 | 23 194 | 2.0 |

Tuning instrument...

TOF Detector optimization started

Optimal Detector voltage: 2400

Setting Detector voltage for both polarities.

Resolution optimization started

Original Values

RO2: 26.83

VS1: -0.20

HST: 0.95

FOR: 5.19

OFS: -34.47

MGV: -973.14

RO2: 27.00

Adjusting ITC to: 4.8

VS1: -0.24

Adjusting ITC to: 7.8

HST: 0.45

OFS: -34.49

MGV: -973.22

FOR: 6.71

MGV: -973.38

Final Values

RO2: 27.00

VS1: -0.24

HST: 0.45

FOR: 6.71

OFS: -34.49

MGV: -973.38

Orignal Aligment Values

Chan 0 Alignment Positive0.00283,-27.509846

Chan 2 Alignment Positive-0.061632,0.723081

Chan 3 Alignment Positive-0.101232,-8.625589

Chan 0 Alignment Negative-0.005767,-26.742754

Chan 2 Alignment Negative-0.022015,1.392134

Chan 3 Alignment Negative-0.031232,-8.062721

Aligning channels

Final Aligment Values

Chan 0 Alignment Positive0.00283,-27.509846

Chan 2 Alignment Positive-0.080699,2.265265

Chan 3 Alignment Positive-0.111728,-8.720874

Chan 0 Alignment Negative-0.005767,-26.742754

Chan 2 Alignment Negative-0.022015,1.392134

Chan 3 Alignment Negative-0.031232,-8.062721

TOF Calibration started

Resolution optimization started

Original Values

VS1: 0.08

VF1: 27.52

HST: -1.15

MGV: -972.37

Adjusting ITC to: 100.0

AAO_AAOEventSampleChangeStatus

Acquisition was aborted.

TOF Resolution optimization cancelled