2012-05-29 at 14:35

Logged in as 9WGC8P1\Administrator

Instrument: Triple TOF 5600

Model #: 1032150/N

Serial #: AY21451105

Instrument Optimization Ver: 3.4172.40

Instrument performance before tuning...

TOF High Resolution

| Mass (Da) | Found At (Da) | Height (cps) | Area | Resolution | Error (ppm) |
| --- | --- | --- | --- | --- | --- |
| 146.1176 | 146.1496 | 2.00E+03 | 8.65E+03 | 22 782 | 219.3 |
| 266.1598 | 266.2189 | 2.48E+03 | 1.27E+04 | 28 254 | 222.1 |
| 315.1623 | 315.2326 | 1.61E+03 | 7.68E+03 | 30 434 | 223.2 |
| 354.2122 | 354.2915 | 1.62E+03 | 7.65E+03 | 33 149 | 223.7 |
| 442.2647 | 442.3639 | 1.96E+03 | 1.00E+04 | 33 447 | 224.3 |
| 609.2807 | 609.4173 | 4.52E+02 | 2.67E+03 | 36 796 | 224.3 |
| 618.3695 | 618.5085 | 2.40E+03 | 1.32E+04 | 35 176 | 224.7 |

Product Ion High Resolution

| Mass (Da) | Found At (Da) | Height (cps) | Area | Resolution | Error (ppm) |
| --- | --- | --- | --- | --- | --- |
| 174.0913 | 174.0986 | 1.59E+04 | 6.53E+04 | 26 124 | 41.7 |
| 195.0652 | 195.0736 | 2.56E+04 | 1.08E+05 | 26 749 | 43.0 |
| 236.1281 | 236.1380 | 4.91E+03 | 2.11E+04 | 28 981 | 41.8 |
| 365.1860 | 365.2018 | 5.74E+03 | 2.77E+04 | 31 566 | 43.2 |
| 397.2122 | 397.2294 | 1.63E+04 | 7.82E+04 | 33 435 | 43.2 |
| 448.1966 | 448.2160 | 9.51E+03 | 4.68E+04 | 34 133 | 43.2 |
| 609.2807 | 609.3069 | 1.27E+04 | 7.48E+04 | 34 615 | 43.1 |

Product Ion High Sensitivity

| Mass (Da) | Found At (Da) | Height (cps) | Area | Resolution | Error (ppm) |
| --- | --- | --- | --- | --- | --- |
| 174.0913 | 174.1298 | 3.88E+04 | 2.58E+05 | 15 639 | 221.0 |
| 195.0652 | 195.1084 | 5.72E+04 | 3.43E+05 | 18 393 | 221.7 |
| 236.1281 | 236.1803 | 1.25E+04 | 1.04E+05 | 14 573 | 221.0 |
| 365.1860 | 365.2674 | 1.35E+04 | 1.30E+05 | 15 258 | 223.0 |
| 397.2122 | 397.3009 | 3.25E+04 | 2.89E+05 | 17 655 | 223.2 |
| 448.1966 | 448.2968 | 2.00E+04 | 2.02E+05 | 16 243 | 223.5 |
| 609.2807 | 609.4172 | 2.51E+04 | 2.89E+05 | 17 060 | 224.1 |

Tuning instrument...

TOF Detector optimization started

Optimal Detector voltage: 2400

Setting Detector voltage for both polarities.

Resolution optimization started

Original Values

RO2: 26.98

VS1: -0.19

HST: 0.43

FOR: 8.19

OFS: -34.59

MGV: -973.46

Adjusting ITC to: 1.2

RO2: 26.98

Adjusting ITC to: 2.4

VS1: -0.22

HST: 0.71

OFS: -34.38

MGV: -972.76

FOR: 7.41

MGV: -973.16

Final Values

RO2: 26.98

VS1: -0.22

HST: 0.71

FOR: 7.41

OFS: -34.38

MGV: -973.16

Orignal Aligment Values

Chan 0 Alignment Positive0.011492,-27.391167

Chan 2 Alignment Positive-0.068638,0.77783

Chan 3 Alignment Positive-0.115739,-8.325597

Chan 0 Alignment Negative-0.005767,-26.742754

Chan 2 Alignment Negative-0.022015,1.392134

Chan 3 Alignment Negative-0.031232,-8.062721

Aligning channels

Final Aligment Values

Chan 0 Alignment Positive0.004910,-26.684872

Chan 2 Alignment Positive-0.076216,0.792709

Chan 3 Alignment Positive-0.134292,-8.271375

Chan 0 Alignment Negative-0.005767,-26.742754

Chan 2 Alignment Negative-0.022015,1.392134

Chan 3 Alignment Negative-0.031232,-8.062721

TOF Calibration started

Resolution optimization started

Original Values

VS1: 0.08

VF1: 27.65

HST: -1.14

MGV: -973.46

Adjusting ITC to: 1.2

VS1: 0.08

VF1: 27.65

VS1: 0.08

VF1: 27.44

MGV: -972.38

HST: 0.66

Final Values

VS1: 0.08

VF1: 27.44

HST: 0.66

MGV: -972.38

TOF Calibration started

TOF Calibration started

Instrument performance after tuning...

TOF High Resolution

| Mass (Da) | Found At (Da) | Height (cps) | Area | Resolution | Error (ppm) |
| --- | --- | --- | --- | --- | --- |
| 146.1176 | 146.1174 | 1.52E+04 | 5.98E+04 | 24 481 | 1.4 |
| 266.1598 | 266.1594 | 1.52E+04 | 6.29E+04 | 32 120 | 1.5 |
| 315.1623 | 315.1622 | 1.53E+04 | 6.85E+04 | 32 203 | 0.3 |
| 354.2122 | 354.2120 | 9.19E+03 | 4.09E+04 | 34 836 | 0.7 |
| 442.2647 | 442.2645 | 1.16E+04 | 5.31E+04 | 37 533 | 0.4 |
| 609.2807 | 609.2804 | 1.05E+04 | 5.40E+04 | 38 904 | 0.5 |
| 618.3695 | 618.3694 | 1.40E+04 | 7.40E+04 | 38 042 | 0.3 |

Product Ion High Resolution

| Mass (Da) | Found At (Da) | Height (cps) | Area | Resolution | Error (ppm) |
| --- | --- | --- | --- | --- | --- |
| 174.0913 | 174.0913 | 5.53E+04 | 2.22E+05 | 26 669 | 0.2 |
| 195.0652 | 195.0652 | 8.62E+04 | 3.16E+05 | 30 983 | 0.2 |
| 236.1281 | 236.1278 | 2.10E+04 | 8.56E+04 | 31 071 | 1.3 |
| 365.1860 | 365.1857 | 2.50E+04 | 1.09E+05 | 35 838 | 0.7 |
| 397.2122 | 397.2125 | 6.16E+04 | 2.65E+05 | 37 825 | 0.7 |
| 448.1966 | 448.1966 | 3.74E+04 | 1.75E+05 | 36 796 | 0.0 |
| 609.2807 | 609.2805 | 5.31E+04 | 2.58E+05 | 40 801 | 0.3 |

Product Ion High Sensitivity

| Mass (Da) | Found At (Da) | Height (cps) | Area | Resolution | Error (ppm) |
| --- | --- | --- | --- | --- | --- |
| 174.0913 | 174.0914 | 7.03E+04 | 3.55E+05 | 20 832 | 0.3 |
| 195.0652 | 195.0651 | 9.20E+04 | 4.12E+05 | 24 844 | 0.3 |
| 236.1281 | 236.1279 | 2.73E+04 | 1.77E+05 | 19 240 | 0.9 |
| 365.1860 | 365.1862 | 2.89E+04 | 2.18E+05 | 20 190 | 0.5 |
| 397.2122 | 397.2123 | 5.95E+04 | 3.88E+05 | 24 336 | 0.2 |
| 448.1966 | 448.1967 | 3.95E+04 | 3.05E+05 | 21 764 | 0.3 |
| 609.2807 | 609.2803 | 4.72E+04 | 3.86E+05 | 23 877 | 0.5 |

**************

Instrument Optimization has modified the instrument settings with the best values it has found. If you are not satisfied with the performance of the instrument, you can restore the settings to a previous state by using the 'File / Restore Instrument Settings' command.

**************

Done