2012-08-01 at 10:03

Logged in as 9WGC8P1\Administrator

Instrument: Triple TOF 5600

Model #: 1032150/N

Serial #: AY21451105

Instrument Optimization Ver: 3.4172.40

Instrument performance before tuning...

TOF High Resolution

| Mass (Da) | Found At (Da) | Height (cps) | Area | Resolution | Error (ppm) |
| --- | --- | --- | --- | --- | --- |
| 144.1030 | 0.0000 | 0.00E+00 | 0.00E+00 | 0 000 | 1000000.0 |
| 264.1453 | 0.0000 | 0.00E+00 | 0.00E+00 | 0 000 | 1000000.0 |
| 352.1977 | 0.0000 | 0.00E+00 | 0.00E+00 | 0 000 | 1000000.0 |
| 403.1122 | 0.0000 | 0.00E+00 | 0.00E+00 | 0 000 | 1000000.0 |
| 440.2501 | 0.0000 | 0.00E+00 | 0.00E+00 | 0 000 | 1000000.0 |
| 616.3550 | 0.0000 | 0.00E+00 | 0.00E+00 | 0 000 | 1000000.0 |
| 792.4598 | 0.0000 | 0.00E+00 | 0.00E+00 | 0 000 | 1000000.0 |

Product Ion High Resolution

| Mass (Da) | Found At (Da) | Height (cps) | Area | Resolution | Error (ppm) |
| --- | --- | --- | --- | --- | --- |
| 93.0344 | 0.0000 | 0.00E+00 | 0.00E+00 | 0 000 | 1000000.0 |
| 125.0067 | 0.0000 | 0.00E+00 | 0.00E+00 | 0 000 | 1000000.0 |
| 158.0611 | 0.0000 | 0.00E+00 | 0.00E+00 | 0 000 | 1000000.0 |
| 277.0983 | 0.0000 | 0.00E+00 | 0.00E+00 | 0 000 | 1000000.0 |
| 403.1122 | 0.0000 | 0.00E+00 | 0.00E+00 | 0 000 | 1000000.0 |

Product Ion High Sensitivity

| Mass (Da) | Found At (Da) | Height (cps) | Area | Resolution | Error (ppm) |
| --- | --- | --- | --- | --- | --- |
| 93.0344 | 0.0000 | 0.00E+00 | 0.00E+00 | 0 000 | 1000000.0 |
| 125.0067 | 0.0000 | 0.00E+00 | 0.00E+00 | 0 000 | 1000000.0 |
| 158.0611 | 0.0000 | 0.00E+00 | 0.00E+00 | 0 000 | 1000000.0 |
| 277.0983 | 0.0000 | 0.00E+00 | 0.00E+00 | 0 000 | 1000000.0 |
| 403.1122 | 0.0000 | 0.00E+00 | 0.00E+00 | 0 000 | 1000000.0 |

Tuning instrument...

TOF Detector optimization started

No signal detected, leaving CEM unchanged

Resolution optimization started

Original Values

RO2: -27.73

VS1: 0.20

HST: -0.23

FOR: -3.70

OFS: -27.78

MGV: 971.81

Adjusting ITC to: 100.0

Low Signal, re-acquiring data

Low Signal, re-acquiring data

Low Signal, re-acquiring data

Low Signal, re-acquiring data

Low Signal, re-acquiring data

Low Signal, re-acquiring data

Low Signal, re-acquiring data

Low Signal, re-acquiring data

Low Signal, re-acquiring data

TOF Resolution optimization cancelled