2012-09-20 at 16:00

Logged in as 9WGC8P1\Administrator

Instrument: Triple TOF 5600

Model #: 1032150/N

Serial #: AY21451105

Instrument Optimization Ver: 3.4172.40

Instrument performance before tuning...

TOF High Resolution

| Mass (Da) | Found At (Da) | Height (cps) | Area | Resolution | Error (ppm) |
| --- | --- | --- | --- | --- | --- |
| 146.1176 | 0.0000 | 0.00E+00 | 0.00E+00 | 0 000 | 1000000.0 |
| 266.1598 | 0.0000 | 0.00E+00 | 0.00E+00 | 0 000 | 1000000.0 |
| 315.1623 | 0.0000 | 0.00E+00 | 0.00E+00 | 0 000 | 1000000.0 |
| 354.2122 | 0.0000 | 0.00E+00 | 0.00E+00 | 0 000 | 1000000.0 |
| 442.2647 | 0.0000 | 0.00E+00 | 0.00E+00 | 0 000 | 1000000.0 |
| 609.2807 | 0.0000 | 0.00E+00 | 0.00E+00 | 0 000 | 1000000.0 |
| 618.3695 | 0.0000 | 0.00E+00 | 0.00E+00 | 0 000 | 1000000.0 |
| 922.0098 | 0.0000 | 0.00E+00 | 0.00E+00 | 0 000 | 1000000.0 |

Product Ion High Resolution

| Mass (Da) | Found At (Da) | Height (cps) | Area | Resolution | Error (ppm) |
| --- | --- | --- | --- | --- | --- |
| 174.0913 | 0.0000 | 0.00E+00 | 0.00E+00 | 0 000 | 1000000.0 |
| 195.0652 | 0.0000 | 0.00E+00 | 0.00E+00 | 0 000 | 1000000.0 |
| 236.1281 | 0.0000 | 0.00E+00 | 0.00E+00 | 0 000 | 1000000.0 |
| 365.1860 | 0.0000 | 0.00E+00 | 0.00E+00 | 0 000 | 1000000.0 |
| 397.2122 | 0.0000 | 0.00E+00 | 0.00E+00 | 0 000 | 1000000.0 |
| 448.1966 | 0.0000 | 0.00E+00 | 0.00E+00 | 0 000 | 1000000.0 |
| 609.2807 | 0.0000 | 0.00E+00 | 0.00E+00 | 0 000 | 1000000.0 |

Product Ion High Sensitivity

| Mass (Da) | Found At (Da) | Height (cps) | Area | Resolution | Error (ppm) |
| --- | --- | --- | --- | --- | --- |
| 174.0913 | 174.0830 | 5.53E+03 | 3.85E+04 | 16 294 | 47.8 |
| 195.0652 | 195.0561 | 8.84E+03 | 5.99E+04 | 17 678 | 46.7 |
| 236.1281 | 236.1167 | 1.71E+03 | 1.37E+04 | 17 059 | 48.1 |
| 365.1860 | 365.1688 | 1.82E+03 | 1.76E+04 | 16 608 | 47.2 |
| 397.2122 | 397.1942 | 5.04E+03 | 4.93E+04 | 17 128 | 45.4 |
| 448.1966 | 448.1760 | 2.78E+03 | 2.94E+04 | 16 364 | 45.9 |
| 609.2807 | 609.2525 | 4.35E+03 | 5.49E+04 | 16 188 | 46.3 |

Tuning instrument...

TOF Detector optimization started

Optimal Detector voltage: 2050

Setting Detector voltage for both polarities.

Resolution optimization started

Original Values

RO2: 27.02

VS1: -0.17

HST: 0.68

FOR: 7.42

OFS: -34.50

MGV: -973.30

Adjusting ITC to: 100.0

Low Signal, re-acquiring data

Adjusting ITC to: 12.9

Adjusting ITC to: 4.8

RO2: 26.57

Adjusting ITC to: 45.0

Adjusting ITC to: 18.7

VS1: -0.18

HST: 0.21

OFS: -36.79

MGV: -972.82

Adjusting ITC to: 12.5

FOR: 3.62

MGV: -972.69

Final Values

RO2: 26.57

VS1: -0.18

HST: 0.21

FOR: 3.62

OFS: -36.79

MGV: -972.69

Orignal Aligment Values

Chan 0 Alignment Positive0.007033,-27.550974

Chan 2 Alignment Positive-0.072949,0.56712

Chan 3 Alignment Positive-0.132905,-7.779977

Chan 0 Alignment Negative0.007697,-26.039093

Chan 2 Alignment Negative-0.033586,0.215768

Chan 3 Alignment Negative-0.064246,-8.003115

Aligning channels

Final Aligment Values

Chan 0 Alignment Positive0.041620,-27.116888

Chan 2 Alignment Positive-0.074799,0.250228

Chan 3 Alignment Positive-0.110708,-8.816854

Chan 0 Alignment Negative0.007697,-26.039093

Chan 2 Alignment Negative-0.033586,0.215768

Chan 3 Alignment Negative-0.064246,-8.003115

TOF Calibration started

Resolution optimization started

Original Values

VS1: 0.08

VF1: 27.90

HST: 0.43

MGV: -972.65

VS1: 0.08

Adjusting ITC to: 6.0

VF1: 27.86

Adjusting ITC to: 11.3

Adjusting ITC to: 7.6

VS1: 0.08

VF1: 27.86

MGV: -973.06

Adjusting ITC to: 4.6

HST: -1.36

Final Values

VS1: 0.08

VF1: 27.86

HST: -1.36

MGV: -972.69

TOF Calibration started

TOF Calibration started

Instrument performance after tuning...

TOF High Resolution

| Mass (Da) | Found At (Da) | Height (cps) | Area | Resolution | Error (ppm) |
| --- | --- | --- | --- | --- | --- |
| 146.1176 | 146.1172 | 3.29E+03 | 1.00E+04 | 30 445 | 2.3 |
| 266.1598 | 266.1594 | 1.42E+03 | 5.26E+03 | 34 763 | 1.4 |
| 315.1623 | 315.1617 | 1.04E+03 | 3.66E+03 | 40 787 | 1.7 |
| 354.2122 | 354.2119 | 6.04E+02 | 2.40E+03 | 37 662 | 1.1 |
| 442.2647 | 442.2643 | 6.45E+02 | 2.80E+03 | 38 251 | 0.8 |
| 609.2807 | 609.2798 | 4.61E+02 | 1.95E+03 | 46 813 | 1.3 |
| 618.3695 | 618.3686 | 1.10E+03 | 4.55E+03 | 47 534 | 1.4 |
| 922.0098 | 922.0081 | 2.00E+01 | 6.20E+01 | 86 299 | 1.8 |

Product Ion High Resolution

| Mass (Da) | Found At (Da) | Height (cps) | Area | Resolution | Error (ppm) |
| --- | --- | --- | --- | --- | --- |
| 174.0913 | 174.0913 | 1.31E+04 | 4.18E+04 | 31 605 | 0.1 |
| 195.0652 | 195.0651 | 3.00E+04 | 9.77E+04 | 32 873 | 0.7 |
| 236.1281 | 236.1278 | 2.03E+03 | 7.50E+03 | 32 570 | 1.4 |
| 365.1860 | 365.1857 | 2.35E+03 | 8.87E+03 | 40 077 | 0.7 |
| 397.2122 | 397.2119 | 1.48E+04 | 5.61E+04 | 41 603 | 0.7 |
| 448.1966 | 448.1963 | 4.89E+03 | 2.00E+04 | 40 258 | 0.8 |
| 609.2807 | 609.2801 | 1.13E+04 | 5.04E+04 | 43 456 | 0.9 |

Product Ion High Sensitivity

| Mass (Da) | Found At (Da) | Height (cps) | Area | Resolution | Error (ppm) |
| --- | --- | --- | --- | --- | --- |
| 174.0913 | 174.0910 | 3.26E+04 | 1.71E+05 | 19 595 | 1.7 |
| 195.0652 | 195.0648 | 5.81E+04 | 2.88E+05 | 21 748 | 2.3 |
| 236.1281 | 236.1277 | 5.56E+03 | 3.72E+04 | 18 099 | 1.7 |
| 365.1860 | 365.1859 | 6.47E+03 | 4.97E+04 | 19 389 | 0.3 |
| 397.2122 | 397.2117 | 3.06E+04 | 2.12E+05 | 22 282 | 1.2 |
| 448.1966 | 448.1964 | 1.25E+04 | 1.00E+05 | 20 473 | 0.6 |
| 609.2807 | 609.2796 | 2.47E+04 | 2.13E+05 | 22 242 | 1.8 |

**************

Instrument Optimization has modified the instrument settings with the best values it has found. If you are not satisfied with the performance of the instrument, you can restore the settings to a previous state by using the 'File / Restore Instrument Settings' command.

**************

Done