2013-01-17 at 11:38

Logged in as 9WGC8P1\Administrator

Instrument: Triple TOF 5600

Model #: 1032150/N

Serial #: AY21451105

Trying to switch to Acquire mode, attempt# 1

Instrument Optimization Ver: 3.4172.40

Instrument performance before tuning...

TOF High Resolution

| Mass (Da) | Found At (Da) | Height (cps) | Area | Resolution | Error (ppm) |
| --- | --- | --- | --- | --- | --- |
| 132.9049 | 132.9049 | 5.33E+03 | 2.05E+04 | 24 508 | 0.0 |
| 829.5393 | 829.5409 | 6.16E+03 | 3.56E+04 | 39 460 | 2.0 |

Product Ion High Resolution

| Mass (Da) | Found At (Da) | Height (cps) | Area | Resolution | Error (ppm) |
| --- | --- | --- | --- | --- | --- |
| 185.1285 | 185.1223 | 1.29E+04 | 4.94E+04 | 29 257 | 33.4 |
| 215.1390 | 215.1318 | 8.16E+03 | 3.23E+04 | 30 443 | 33.4 |
| 298.2125 | 298.2029 | 3.21E+04 | 1.30E+05 | 34 387 | 32.1 |
| 381.2496 | 381.2375 | 3.34E+04 | 1.40E+05 | 38 318 | 31.7 |
| 494.3337 | 494.3177 | 5.10E+04 | 2.27E+05 | 40 372 | 32.4 |
| 607.4178 | 607.3978 | 3.88E+04 | 1.87E+05 | 41 017 | 32.9 |
| 625.4283 | 625.4075 | 1.92E+04 | 9.85E+04 | 39 537 | 33.2 |
| 724.4967 | 724.4724 | 1.39E+04 | 7.49E+04 | 40 028 | 33.6 |
| 811.5288 | 811.5013 | 2.21E+03 | 1.25E+04 | 40 344 | 33.9 |

Tuning instrument...

TOF Detector optimization started

Optimal Detector voltage: 2150

Setting Detector voltage for both polarities.

Resolution optimization started

Original Values

RO2: 26.81

VS1: -0.24

HST: 0.20

FOR: 6.52

OFS: -36.86

MGV: -973.42

Adjusting ITC to: 2.0

RO2: 26.79

Adjusting ITC to: 5.4

VS1: -0.22

HST: 0.00

OFS: -39.55

MGV: -972.96

FOR: 7.41

Adjusting ITC to: 8.3

MGV: -973.01

Final Values

RO2: 26.79

VS1: -0.22

HST: 0.00

FOR: 7.41

OFS: -39.55

MGV: -973.01

Orignal Aligment Values

Chan 0 Alignment Positive0.0259,-26.643952

Chan 2 Alignment Positive-0.080666,0.812476

Chan 3 Alignment Positive-0.138278,-7.868763

Chan 0 Alignment Negative0.020698,-25.849047

Chan 2 Alignment Negative-0.035159,0.027415

Chan 3 Alignment Negative-0.049992,-8.294935

Aligning channels

Final Aligment Values

Chan 0 Alignment Positive0.033505,-28.048640

Chan 2 Alignment Positive-0.071757,0.322801

Chan 3 Alignment Positive-0.118723,-8.703154

Chan 0 Alignment Negative0.020698,-25.849047

Chan 2 Alignment Negative-0.035159,0.027415

Chan 3 Alignment Negative-0.049992,-8.294935

TOF Calibration started

TOF Calibration started

Instrument performance after tuning...

TOF High Resolution

| Mass (Da) | Found At (Da) | Height (cps) | Area | Resolution | Error (ppm) |
| --- | --- | --- | --- | --- | --- |
| 132.9049 | 132.9046 | 4.49E+03 | 1.59E+04 | 25 212 | 2.0 |
| 829.5393 | 829.5385 | 4.20E+03 | 2.68E+04 | 36 208 | 1.0 |

Product Ion High Resolution

| Mass (Da) | Found At (Da) | Height (cps) | Area | Resolution | Error (ppm) |
| --- | --- | --- | --- | --- | --- |
| 185.1285 | 185.1285 | 1.22E+04 | 4.90E+04 | 26 903 | 0.0 |
| 215.1390 | 215.1391 | 7.55E+03 | 3.19E+04 | 27 546 | 0.4 |
| 298.2125 | 298.2130 | 2.76E+04 | 1.24E+05 | 30 452 | 1.8 |
| 381.2496 | 381.2503 | 2.96E+04 | 1.33E+05 | 34 390 | 1.8 |
| 494.3337 | 494.3347 | 4.53E+04 | 2.15E+05 | 37 444 | 1.9 |
| 607.4178 | 607.4185 | 3.14E+04 | 1.68E+05 | 36 590 | 1.2 |
| 625.4283 | 625.4288 | 1.47E+04 | 8.31E+04 | 35 393 | 0.8 |
| 724.4967 | 724.4976 | 1.03E+04 | 5.95E+04 | 36 709 | 1.2 |
| 811.5288 | 811.5292 | 1.49E+03 | 9.36E+03 | 35 073 | 0.5 |

**************

Instrument Optimization has modified the instrument settings with the best values it has found. If you are not satisfied with the performance of the instrument, you can restore the settings to a previous state by using the 'File / Restore Instrument Settings' command.

**************

Done