2013-01-17 at 11:52

Logged in as 9WGC8P1\Administrator

Instrument: Triple TOF 5600

Model #: 1032150/N

Serial #: AY21451105

Instrument Optimization Ver: 3.4172.40

Instrument performance before tuning...

Product Ion High Sensitivity

| Mass (Da) | Found At (Da) | Height (cps) | Area | Resolution | Error (ppm) |
| --- | --- | --- | --- | --- | --- |
| 185.1285 | 185.1016 | 9.67E+03 | 7.02E+04 | 14 710 | 145.2 |
| 215.1390 | 215.1078 | 6.09E+03 | 4.59E+04 | 15 076 | 145.1 |
| 298.2125 | 298.1694 | 2.02E+04 | 1.59E+05 | 17 170 | 144.5 |
| 381.2496 | 381.1945 | 1.95E+04 | 1.67E+05 | 17 902 | 144.5 |
| 494.3337 | 494.2615 | 2.86E+04 | 2.42E+05 | 20 750 | 146.0 |
| 607.4178 | 607.3290 | 2.22E+04 | 2.15E+05 | 20 034 | 146.2 |
| 625.4283 | 625.3373 | 1.14E+04 | 1.28E+05 | 17 334 | 145.4 |
| 724.4967 | 724.3913 | 8.16E+03 | 1.01E+05 | 16 727 | 145.5 |
| 811.5288 | 811.4107 | 1.25E+03 | 1.77E+04 | 15 050 | 145.5 |

Tuning instrument...

Resolution optimization started

Original Values

VS1: 0.08

VF1: 27.88

HST: -1.15

MGV: -973.42

Adjusting ITC to: 4.9

VS1: 0.11

Adjusting ITC to: 2.4

VF1: 27.87

VS1: 0.08

VF1: 27.87

MGV: -973.17

HST: -1.14

Final Values

VS1: 0.08

VF1: 27.87

HST: -1.14

MGV: -973.01

TOF Calibration started

Instrument performance after tuning...

Product Ion High Sensitivity

| Mass (Da) | Found At (Da) | Height (cps) | Area | Resolution | Error (ppm) |
| --- | --- | --- | --- | --- | --- |
| 185.1285 | 185.1284 | 1.78E+04 | 1.25E+05 | 15 189 | 0.7 |
| 215.1390 | 215.1388 | 1.13E+04 | 8.34E+04 | 15 338 | 0.9 |
| 298.2125 | 298.2123 | 3.58E+04 | 2.60E+05 | 18 591 | 0.8 |
| 381.2496 | 381.2496 | 3.41E+04 | 2.72E+05 | 19 206 | 0.1 |
| 494.3337 | 494.3329 | 4.60E+04 | 3.53E+05 | 22 918 | 1.7 |
| 607.4178 | 607.4169 | 3.50E+04 | 3.20E+05 | 21 168 | 1.5 |
| 625.4283 | 625.4281 | 1.83E+04 | 2.03E+05 | 17 746 | 0.3 |
| 724.4967 | 724.4961 | 1.30E+04 | 1.58E+05 | 17 045 | 0.8 |
| 811.5288 | 811.5279 | 2.00E+03 | 2.86E+04 | 15 128 | 1.1 |

**************

Instrument Optimization has modified the instrument settings with the best values it has found. If you are not satisfied with the performance of the instrument, you can restore the settings to a previous state by using the 'File / Restore Instrument Settings' command.

**************

Done