2013-08-08 at 17:06

Logged in as 9WGC8P1\Administrator

Instrument: Triple TOF 5600

Model #: 1032150/N

Serial #: AY21451105

Instrument Optimization Ver: 3.4172.40

Instrument performance before tuning...

TOF High Resolution

| Mass (Da) | Found At (Da) | Height (cps) | Area | Resolution | Error (ppm) |
| --- | --- | --- | --- | --- | --- |
| 146.1176 | 146.1156 | 4.27E+02 | 1.34E+03 | 32 915 | 13.7 |
| 266.1598 | 266.1563 | 2.33E+02 | 9.32E+02 | 32 561 | 13.2 |
| 315.1623 | 315.1581 | 2.01E+02 | 8.84E+02 | 32 128 | 13.2 |
| 354.2122 | 354.0758 | 1.41E+02 | 5.62E+02 | 37 928 | 385.3 |
| 442.2647 | 442.2590 | 1.47E+02 | 6.79E+02 | 37 783 | 12.8 |
| 609.2807 | 609.1011 | 1.38E+02 | 6.95E+02 | 40 647 | 294.7 |
| 618.3695 | 618.1883 | 2.08E+02 | 9.56E+02 | 44 458 | 293.0 |
| 922.0098 | 921.7870 | 3.20E+01 | 1.61E+02 | 43 015 | 241.6 |

Product Ion High Resolution

| Mass (Da) | Found At (Da) | Height (cps) | Area | Resolution | Error (ppm) |
| --- | --- | --- | --- | --- | --- |
| 174.0913 | 174.1073 | 4.36E+03 | 1.55E+04 | 30 887 | 92.0 |
| 195.0652 | 195.0830 | 7.87E+03 | 2.71E+04 | 33 098 | 91.5 |
| 236.1281 | 236.1494 | 1.42E+03 | 5.26E+03 | 34 393 | 90.4 |
| 365.1860 | 365.2193 | 1.56E+03 | 6.62E+03 | 37 148 | 91.1 |
| 397.2122 | 397.1096 | 4.93E+03 | 2.06E+04 | 38 859 | 258.4 |
| 448.1966 | 448.0899 | 2.54E+03 | 1.14E+04 | 37 911 | 238.1 |
| 609.2807 | 609.3355 | 4.04E+03 | 2.07E+04 | 38 437 | 90.0 |

Product Ion High Sensitivity

| Mass (Da) | Found At (Da) | Height (cps) | Area | Resolution | Error (ppm) |
| --- | --- | --- | --- | --- | --- |
| 174.0913 | 173.9936 | 4.00E+00 | 3.00E+00 | 0 001 | 561.3 |
| 195.0652 | 195.0608 | 4.00E+00 | 7.00E+00 | 53 963 | 22.7 |
| 236.1281 | 195.0608 | 4.00E+00 | 0.00E+00 | 53 963 | 173919.7 |
| 365.1860 | 195.0608 | 4.00E+00 | 0.00E+00 | 53 963 | 465859.1 |
| 397.2122 | 195.0608 | 4.00E+00 | 0.00E+00 | 53 963 | 508925.5 |
| 448.1966 | 195.0608 | 4.00E+00 | 0.00E+00 | 53 963 | 564787.5 |
| 609.2807 | 195.0608 | 4.00E+00 | 0.00E+00 | 53 963 | 679850.7 |

Tuning instrument...

TOF Detector optimization started

Optimal Detector voltage: 2150

Setting Detector voltage for both polarities.

Resolution optimization started

Original Values

RO2: 26.98

VS1: -0.20

HST: -0.19

FOR: 9.13

OFS: -36.95

MGV: -972.76

Adjusting ITC to: 6.6

Adjusting ITC to: 13.1

Adjusting ITC to: 100.0

Low Signal, re-acquiring data

Low Signal, re-acquiring data

Low Signal, re-acquiring data

Adjusting ITC to: 24.4

Adjusting ITC to: 56.0

Adjusting ITC to: 100.0

Low Signal, re-acquiring data

TOF Resolution optimization cancelled