2013-10-01 at 09:22

Logged in as 9WGC8P1\Administrator

Instrument: Triple TOF 5600

Model #: 1032150/N

Serial #: AY21451105

Instrument Optimization Ver: 3.4172.40

Instrument performance before tuning...

TOF High Resolution

| Mass (Da) | Found At (Da) | Height (cps) | Area | Resolution | Error (ppm) |
| --- | --- | --- | --- | --- | --- |
| 146.1176 | 146.1168 | 5.60E+03 | 1.99E+04 | 27 444 | 5.4 |
| 266.1598 | 266.1584 | 5.13E+03 | 2.19E+04 | 30 589 | 5.2 |
| 315.1623 | 315.1606 | 1.85E+03 | 8.48E+03 | 31 202 | 5.2 |
| 354.2122 | 354.2105 | 3.04E+03 | 1.38E+04 | 33 451 | 4.8 |
| 442.2647 | 442.2627 | 3.22E+03 | 1.58E+04 | 34 156 | 4.6 |
| 609.2807 | 609.2776 | 9.68E+02 | 5.44E+03 | 36 726 | 5.1 |
| 618.3695 | 618.3668 | 4.19E+03 | 2.39E+04 | 34 117 | 4.4 |
| 922.0098 | 922.0054 | 3.07E+02 | 2.09E+03 | 34 783 | 4.7 |

Product Ion High Resolution

| Mass (Da) | Found At (Da) | Height (cps) | Area | Resolution | Error (ppm) |
| --- | --- | --- | --- | --- | --- |
| 174.0913 | 174.0959 | 1.50E+03 | 5.53E+03 | 27 937 | 26.6 |
| 195.0652 | 195.0703 | 2.65E+03 | 1.01E+04 | 28 725 | 26.3 |
| 236.1281 | 236.1343 | 3.99E+02 | 1.76E+03 | 28 203 | 26.5 |
| 365.1860 | 365.1961 | 4.77E+02 | 2.24E+03 | 34 127 | 27.5 |
| 397.2122 | 397.2233 | 1.53E+03 | 7.27E+03 | 33 517 | 27.9 |
| 448.1966 | 448.2091 | 7.57E+02 | 3.83E+03 | 33 791 | 27.8 |
| 609.2807 | 609.2972 | 1.27E+03 | 7.30E+03 | 34 259 | 27.2 |

Product Ion High Sensitivity

| Mass (Da) | Found At (Da) | Height (cps) | Area | Resolution | Error (ppm) |
| --- | --- | --- | --- | --- | --- |
| 174.0913 | 174.0911 | 2.00E+03 | 1.21E+04 | 17 014 | 1.4 |
| 195.0652 | 195.0650 | 3.52E+03 | 2.12E+04 | 17 737 | 1.2 |
| 236.1281 | 236.1277 | 5.67E+02 | 3.61E+03 | 18 572 | 1.6 |
| 365.1860 | 365.1859 | 5.91E+02 | 4.70E+03 | 18 677 | 0.3 |
| 397.2122 | 397.2121 | 1.73E+03 | 1.51E+04 | 17 517 | 0.2 |
| 448.1966 | 448.1963 | 9.21E+02 | 8.22E+03 | 17 965 | 0.7 |
| 609.2807 | 609.2797 | 1.39E+03 | 1.54E+04 | 17 065 | 1.6 |

Tuning instrument...

TOF Detector optimization started

Optimal Detector voltage: 2100

Setting Detector voltage for both polarities.

Resolution optimization started

Original Values

RO2: 27.01

VS1: -0.20

HST: -0.19

FOR: 8.97

OFS: -37.27

MGV: -972.68

Adjusting ITC to: 100.0

Low Signal, re-acquiring data

Low Signal, re-acquiring data

Low Signal, re-acquiring data

Low Signal, re-acquiring data

Adjusting ITC to: 54.5

Adjusting ITC to: 13.8

Adjusting ITC to: 10.0

Adjusting ITC to: 100.0

Adjusting ITC to: 53.8

TOF Resolution optimization cancelled