2013-10-01 at 09:36

Logged in as 9WGC8P1\Administrator

Instrument: Triple TOF 5600

Model #: 1032150/N

Serial #: AY21451105

Instrument Optimization Ver: 3.4172.40

Instrument performance before tuning...

TOF High Resolution

| Mass (Da) | Found At (Da) | Height (cps) | Area | Resolution | Error (ppm) |
| --- | --- | --- | --- | --- | --- |
| 146.1176 | 146.1162 | 8.24E+03 | 2.79E+04 | 28 122 | 9.2 |
| 266.1598 | 266.1573 | 9.44E+03 | 3.85E+04 | 31 657 | 9.3 |
| 315.1623 | 315.1588 | 7.00E+01 | 3.25E+02 | 30 647 | 10.8 |
| 354.2122 | 354.2094 | 8.17E+03 | 3.83E+04 | 31 803 | 7.9 |
| 442.2647 | 442.2610 | 5.88E+03 | 2.84E+04 | 34 761 | 8.4 |
| 609.2807 | 609.2807 | 5.00E+00 | 1.40E+01 | 139 881 | 0.1 |
| 618.3695 | 618.3641 | 6.31E+03 | 3.54E+04 | 34 836 | 8.8 |
| 922.0098 | 618.3641 | 6.31E+03 | 0.00E+00 | 34 836 | 329330.2 |

Product Ion High Resolution

| Mass (Da) | Found At (Da) | Height (cps) | Area | Resolution | Error (ppm) |
| --- | --- | --- | --- | --- | --- |
| 174.0913 | 174.0951 | 3.20E+01 | 1.04E+02 | 35 222 | 22.0 |
| 195.0652 | 195.0696 | 7.80E+01 | 3.03E+02 | 27 396 | 22.3 |
| 236.1281 | 236.1317 | 1.50E+01 | 3.10E+01 | 61 469 | 15.3 |
| 365.1860 | 365.1938 | 1.30E+01 | 3.00E+01 | 67 684 | 21.3 |
| 397.2122 | 397.2216 | 3.50E+01 | 1.67E+02 | 27 887 | 23.8 |
| 448.1966 | 448.2064 | 1.90E+01 | 7.80E+01 | 46 143 | 21.9 |
| 609.2807 | 609.3729 | 4.50E+01 | 2.96E+02 | 30 001 | 151.4 |

Product Ion High Sensitivity

| Mass (Da) | Found At (Da) | Height (cps) | Area | Resolution | Error (ppm) |
| --- | --- | --- | --- | --- | --- |
| 174.0913 | 174.0901 | 9.60E+01 | 5.72E+02 | 16 532 | 6.7 |
| 195.0652 | 195.0638 | 2.01E+02 | 1.28E+03 | 16 210 | 7.0 |
| 236.1281 | 236.1265 | 2.60E+01 | 1.70E+02 | 20 410 | 7.0 |
| 365.1860 | 365.1855 | 2.60E+01 | 2.08E+02 | 20 137 | 1.4 |
| 397.2122 | 397.2096 | 8.40E+01 | 7.45E+02 | 16 749 | 6.6 |
| 448.1966 | 448.1952 | 4.80E+01 | 3.97E+02 | 17 945 | 3.2 |
| 609.2807 | 609.3522 | 7.80E+01 | 9.36E+02 | 14 051 | 117.4 |

Tuning instrument...

TOF Detector optimization started

Optimal Detector voltage: 2150

Setting Detector voltage for both polarities.

Resolution optimization started

Original Values

RO2: 27.01

VS1: -0.20

HST: -0.19

FOR: 8.97

OFS: -37.27

MGV: -972.68

Adjusting ITC to: 100.0

Low Signal, re-acquiring data

Low Signal, re-acquiring data

Low Signal, re-acquiring data

Low Signal, re-acquiring data

AAO_AAOEventSampleChangeStatus

Acquisition was aborted.

TOF Resolution optimization cancelled