2013-10-01 at 10:29

Logged in as 9WGC8P1\Administrator

Instrument: Triple TOF 5600

Model #: 1032150/N

Serial #: AY21451105

Instrument Optimization Ver: 3.4172.40

Instrument performance before tuning...

TOF High Resolution

| Mass (Da) | Found At (Da) | Height (cps) | Area | Resolution | Error (ppm) |
| --- | --- | --- | --- | --- | --- |
| 144.1030 | 144.0990 | 2.00E+00 | 5.00E+00 | 28 343 | 27.8 |
| 264.1453 | 265.1419 | 1.80E+01 | 1.08E+02 | 23 516 | 3772.9 |
| 352.1977 | 352.1861 | 7.00E+00 | 5.00E+01 | 16 616 | 32.9 |
| 403.1122 | 403.0988 | 2.00E+00 | 7.00E+00 | 35 553 | 33.2 |
| 440.2501 | 440.2364 | 1.60E+01 | 9.30E+01 | 30 331 | 31.1 |
| 616.3550 | 616.3377 | 2.70E+01 | 2.44E+02 | 21 981 | 28.1 |
| 792.4598 | 616.3377 | 2.70E+01 | 0.00E+00 | 21 981 | 222247.4 |

Product Ion High Resolution

| Mass (Da) | Found At (Da) | Height (cps) | Area | Resolution | Error (ppm) |
| --- | --- | --- | --- | --- | --- |
| 93.0344 | 0.0000 | 0.00E+00 | 0.00E+00 | 0 000 | 1000000.0 |
| 125.0067 | 125.0168 | 4.10E+01 | 1.96E+02 | 19 263 | 80.6 |
| 158.0611 | 158.0751 | 6.00E+00 | 2.70E+01 | 20 239 | 88.2 |
| 277.0983 | 277.1228 | 2.20E+01 | 1.43E+02 | 19 651 | 88.3 |
| 403.1122 | 403.2316 | 1.00E+01 | 6.80E+01 | 28 445 | 296.2 |

Product Ion High Sensitivity

| Mass (Da) | Found At (Da) | Height (cps) | Area | Resolution | Error (ppm) |
| --- | --- | --- | --- | --- | --- |
| 93.0344 | 0.0000 | 0.00E+00 | 0.00E+00 | 0 000 | 1000000.0 |
| 125.0067 | 125.0022 | 1.00E+00 | 1.00E+00 | 0 001 | 36.4 |
| 158.0611 | 125.0022 | 1.00E+00 | 0.00E+00 | 0 001 | 209153.2 |
| 277.0983 | 277.0916 | 1.00E+00 | 2.00E+00 | 58 955 | 24.1 |
| 403.1122 | 403.1111 | 1.00E+00 | 1.00E+00 | 142 215 | 2.6 |

Tuning instrument...

TOF Detector optimization started

Optimal Detector voltage: 2350

Setting Detector voltage for both polarities.

Resolution optimization started

Original Values

RO2: -26.93

VS1: 0.33

HST: 0.00

FOR: -7.52

OFS: -31.59

MGV: 971.80

Adjusting ITC to: 100.0

Low Signal, re-acquiring data

RO2: -29.50

Low Signal, re-acquiring data

Low Signal, re-acquiring data

Low Signal, re-acquiring data

Low Signal, re-acquiring data

Low Signal, re-acquiring data

Low Signal, re-acquiring data

Low Signal, re-acquiring data

Low Signal, re-acquiring data

Low Signal, re-acquiring data

Low Signal, re-acquiring data

TOF Resolution optimization cancelled