2013-10-01 at 10:51

Logged in as 9WGC8P1\Administrator

Instrument: Triple TOF 5600

Model #: 1032150/N

Serial #: AY21451105

Instrument Optimization Ver: 3.4172.40

Instrument performance before tuning...

TOF High Resolution

| Mass (Da) | Found At (Da) | Height (cps) | Area | Resolution | Error (ppm) |
| --- | --- | --- | --- | --- | --- |
| 144.1030 | 144.0977 | 4.00E+00 | 1.20E+01 | 23 190 | 36.7 |
| 264.1453 | 264.1341 | 1.70E+01 | 7.40E+01 | 34 109 | 42.3 |
| 352.1977 | 352.1847 | 1.90E+01 | 1.06E+02 | 25 669 | 37.0 |
| 403.1122 | 403.0966 | 2.01E+02 | 1.44E+03 | 22 677 | 38.7 |
| 440.2501 | 440.2318 | 1.50E+01 | 8.80E+01 | 32 134 | 41.7 |
| 616.3550 | 616.3310 | 1.90E+01 | 1.47E+02 | 54 042 | 39.0 |
| 792.4598 | 792.4285 | 1.50E+01 | 1.43E+02 | 26 842 | 39.6 |

Product Ion High Resolution

| Mass (Da) | Found At (Da) | Height (cps) | Area | Resolution | Error (ppm) |
| --- | --- | --- | --- | --- | --- |
| 93.0344 | 93.0419 | 3.57E+02 | 1.48E+03 | 19 491 | 80.5 |
| 125.0067 | 125.0156 | 8.61E+03 | 3.80E+04 | 20 827 | 71.4 |
| 158.0611 | 158.0720 | 1.07E+03 | 4.87E+03 | 22 175 | 68.6 |
| 277.0983 | 277.1176 | 5.01E+03 | 3.04E+04 | 22 009 | 69.7 |
| 403.1122 | 403.1408 | 4.17E+02 | 3.00E+03 | 23 456 | 71.1 |

Product Ion High Sensitivity

| Mass (Da) | Found At (Da) | Height (cps) | Area | Resolution | Error (ppm) |
| --- | --- | --- | --- | --- | --- |
| 93.0344 | 93.0324 | 7.28E+02 | 4.02E+03 | 14 122 | 21.8 |
| 125.0067 | 125.0027 | 1.81E+04 | 9.59E+04 | 16 437 | 31.9 |
| 158.0611 | 158.0557 | 2.27E+03 | 1.35E+04 | 16 569 | 34.2 |
| 277.0983 | 277.0897 | 9.78E+03 | 7.46E+04 | 17 216 | 31.2 |
| 403.1122 | 403.1005 | 8.04E+02 | 8.09E+03 | 15 915 | 29.0 |

Tuning instrument...

TOF Detector optimization started

Optimal Detector voltage: 2150

Setting Detector voltage for both polarities.

Resolution optimization started

Original Values

RO2: -26.93

VS1: 0.33

HST: 0.00

FOR: -7.52

OFS: -31.59

MGV: 971.80

RO2: -27.25

Low Signal, re-acquiring data

Low Signal, re-acquiring data

Low Signal, re-acquiring data

Low Signal, re-acquiring data

Low Signal, re-acquiring data

Low Signal, re-acquiring data

Low Signal, re-acquiring data

Low Signal, re-acquiring data

Low Signal, re-acquiring data

Low Signal, re-acquiring data

TOF Resolution optimization cancelled