2013-10-01 at 11:04

Logged in as 9WGC8P1\Administrator

Instrument: Triple TOF 5600

Model #: 1032150/N

Serial #: AY21451105

Instrument Optimization Ver: 3.4172.40

Instrument performance before tuning...

TOF High Resolution

| Mass (Da) | Found At (Da) | Height (cps) | Area | Resolution | Error (ppm) |
| --- | --- | --- | --- | --- | --- |
| 144.1030 | 143.1058 | 3.00E+00 | 1.00E+00 | 0 001 | 6920.3 |
| 264.1453 | 264.1318 | 5.00E+00 | 1.90E+01 | 23 024 | 51.0 |
| 352.1977 | 352.1851 | 6.00E+00 | 3.20E+01 | 20 141 | 35.7 |
| 403.1122 | 403.0967 | 7.80E+01 | 5.41E+02 | 20 988 | 38.4 |
| 440.2501 | 440.2320 | 8.00E+00 | 2.60E+01 | 44 746 | 41.2 |
| 616.3550 | 616.3307 | 1.00E+01 | 5.60E+01 | 63 945 | 39.4 |
| 792.4598 | 792.4289 | 6.00E+00 | 2.20E+01 | 46 916 | 39.1 |

Product Ion High Resolution

| Mass (Da) | Found At (Da) | Height (cps) | Area | Resolution | Error (ppm) |
| --- | --- | --- | --- | --- | --- |
| 93.0344 | 93.0423 | 1.05E+02 | 4.59E+02 | 20 662 | 85.0 |
| 125.0067 | 125.0158 | 2.87E+03 | 1.22E+04 | 21 409 | 72.7 |
| 158.0611 | 158.0722 | 3.15E+02 | 1.49E+03 | 21 110 | 70.1 |
| 277.0983 | 277.1181 | 1.44E+03 | 9.31E+03 | 20 817 | 71.3 |
| 403.1122 | 403.1412 | 1.36E+02 | 9.84E+02 | 23 238 | 72.1 |

Product Ion High Sensitivity

| Mass (Da) | Found At (Da) | Height (cps) | Area | Resolution | Error (ppm) |
| --- | --- | --- | --- | --- | --- |
| 93.0344 | 93.0325 | 2.28E+02 | 1.25E+03 | 13 723 | 20.2 |
| 125.0067 | 125.0030 | 5.84E+03 | 3.21E+04 | 15 877 | 29.9 |
| 158.0611 | 158.0561 | 7.05E+02 | 4.12E+03 | 17 043 | 31.8 |
| 277.0983 | 277.0901 | 3.15E+03 | 2.45E+04 | 16 544 | 29.6 |
| 403.1122 | 403.1011 | 2.55E+02 | 2.41E+03 | 17 074 | 27.5 |

Tuning instrument...

TOF Detector optimization started

Optimal Detector voltage: 2150

Setting Detector voltage for both polarities.

Resolution optimization started

Original Values

RO2: -26.93

VS1: 0.33

HST: 0.00

FOR: -7.52

OFS: -31.59

MGV: 971.80

Low Signal, re-acquiring data

Low Signal, re-acquiring data

Low Signal, re-acquiring data

Low Signal, re-acquiring data

Low Signal, re-acquiring data

Low Signal, re-acquiring data

Low Signal, re-acquiring data

Low Signal, re-acquiring data

Low Signal, re-acquiring data

Low Signal, re-acquiring data

TOF Resolution optimization cancelled