2013-10-04 at 11:01

Logged in as 9WGC8P1\Administrator

Instrument: Triple TOF 5600

Model #: 1032150/N

Serial #: AY21451105

Instrument Optimization Ver: 3.4172.40

Instrument performance before tuning...

TOF High Resolution

| Mass (Da) | Found At (Da) | Height (cps) | Area | Resolution | Error (ppm) |
| --- | --- | --- | --- | --- | --- |
| 132.9049 | 132.8887 | 2.88E+03 | 1.69E+04 | 14 772 | 122.0 |
| 829.5393 | 829.4370 | 3.24E+03 | 4.42E+04 | 14 817 | 123.4 |

Tuning instrument...

TOF Detector optimization started

Optimal Detector voltage: 2200

Setting Detector voltage for both polarities.

Orignal Aligment Values

Chan 0 Alignment Positive0.033224,-26.976896

Chan 2 Alignment Positive-0.091199,-0.053661

Chan 3 Alignment Positive-0.153907,-9.096252

Chan 0 Alignment Negative0.083747,-26.55434

Chan 2 Alignment Negative-0.032269,-0.293664

Chan 3 Alignment Negative-0.018776,-9.24287

Aligning channels

Final Aligment Values

Chan 0 Alignment Positive-0.004464,-30.416066

Chan 2 Alignment Positive-0.000575,-1.112115

Chan 3 Alignment Positive0.048158,-10.138378

Chan 0 Alignment Negative0.083747,-26.55434

Chan 2 Alignment Negative-0.032269,-0.293664

Chan 3 Alignment Negative-0.018776,-9.24287

Instrument performance after tuning...

TOF High Resolution

| Mass (Da) | Found At (Da) | Height (cps) | Area | Resolution | Error (ppm) |
| --- | --- | --- | --- | --- | --- |
| 132.9049 | 132.8866 | 6.96E+03 | 2.68E+04 | 23 775 | 137.9 |
| 829.5393 | 829.4227 | 7.77E+03 | 6.70E+04 | 25 807 | 140.6 |

**************

Instrument Optimization has modified the instrument settings with the best values it has found. If you are not satisfied with the performance of the instrument, you can restore the settings to a previous state by using the 'File / Restore Instrument Settings' command.

**************

Done